Pulse Amplitude Modulation (PAM) fluorometry

- Device: Phyto-Pam-II

- Method: per manufacturer’s instructions

- Results: Maximum quantum yield (Fv/Fm) (photosynthetic efficiency)

- To execute this procedure successfully, we will need the following materials/supplies:

  - Gloves (S, M, L)

  - Sample

  - Cuvettes

  - RO water

  - 70% EtOH

  - Syringes (6 mL)

PAM Method

Preparation and Cleaning

  - Turn PAM ‘ON’

  - Open Phyto-Winn

  - Clean cuvettes

    - 1x EtoH

    - 3x RO

    - 3x SFSW

Running Sample

  - Homogenize sample

  - Pour (or transfer with pipette) sample into cuvette

  - Click ‘Gain’ and write down custody number (Date turns blue)

  - In a second cuvette, use a syringe to extract 5 mLs of sample, attach 0.22 µM filter, and transfer to cuvette as filtrate

  - Wait for PAR sensor to turn white, check Zoff box

  - Once Zoff is complete, return sample (used for gain - unfiltered) to PAM

  - Wait for PAR sensor to turn white, click ‘SAT-Pulse’ and record mean YII

  - Light Curve

    - Turn ‘ON’ the “AL” light

    - Wait for PAR sensor to turn white, press ‘Start’ (10 steps: 18 s each)

    - Once complete, press ‘Fit’, then ‘Calculator’

    - Click ‘export’, and record the reference # between the parenthesis “()”

  - Final Report

    - Select ‘report’ tab at the bottom, and switch to ‘View Mode’

    - Select ALL parameters (listed on the right >>>)

    - Click (last record) (shown on the right >>>>)

    - Click ‘Export’ (at top)

    - Switch back to ‘Measure Mode’

  - Fast Kinetics

    - Select ‘fast kinetics’ tab on bottom

    - Click ‘Averaging’, Wait for PAR sensor to turn white, then press ‘Start’

      - Repeat 2 more times (total 3 times, waiting for PAR sensor each time)

    - Once complete, press ‘Calc’ at the top

    - When next screen appears, click ‘Start fit’, then ‘Export’ file

- If you are running more samples within the same treatment, turn ‘OFF’ the “AL” button, then select ‘Channels’ tab to return to reading YII

- If starting a new treatment completely, close the application and reopen

References:

- PHYTO-PAM-II